# Selective Multipathology Chest X-Ray Classification via Rejection Mechanisms

Yehudit Aperstein1,*, Amit Tzahar1, Alon Gottlib1, Tal Verber1, Ravit Shagan Damti1, Alexander Apartsin2

1 Intelligent Systems, Afeka Academic College of Engineering, Tel Aviv 69988, Israel
2 School of Computer Science, Faculty of Sciences, HIT-Holon Institute of Technology, Holon 58102, Israel
* Corresponding Author: apersteiny@afeka.ac.il

Abstract

Overconfidence in deep learning models poses a significant risk in high-stakes medical imaging tasks, particularly in multi-label classification of chest X-rays, where multiple co-occurring pathologies must be detected simultaneously. This study introduces an uncertainty-aware framework for chest X-ray diagnosis based on a DenseNet-121 backbone, enhanced with an entropy-based selective prediction mechanism that abstains, on a per-pathology basis, from predictions whose binary entropy exceeds a calibrated threshold, deferring those findings to clinical experts. A quantile-based calibration procedure tunes rejection thresholds using either global or class-specific strategies. On NIH ChestX-ray14, MIMIC-CXR, and PadChest, selective prediction is evaluated with risk-coverage analysis and balanced accuracy under non-parametric bootstrap confidence intervals. At a calibrated operating point, abstention improves balanced accuracy for all four target pathologies on NIH ChestX-ray14, with gains of 0.07 to 0.09 at a coverage of 0.70 (95% confidence intervals excluding zero). Every target pathology is also supported on MIMIC-CXR and PadChest, on an independently trained model evaluated cross-dataset on a source it never saw, and under leave-one-dataset-out domain shift in both transfer directions. A key methodological factor is the uncertainty axis: computing predictive entropy on the model's raw probability, rather than the operating-point-normalized output that pretrained backbones emit by default, is necessary for the selective signal to transfer consistently across pathologies and datasets. These results support the integration of selective prediction into AI-assisted diagnostic workflows, providing a practical step toward safer, uncertainty-aware deployment of deep learning in clinical settings.

Keywords: Chest X-ray classification; Rejection mechanism; Multi-label diagnosis; Selective prediction; Uncertainty estimation; Risk-coverage analysis; Balanced accuracy; Domain shift

## 1. Introduction

Automating medical diagnosis with deep learning has shown great potential, particularly in medical imaging domains such as chest X-ray analysis. Convolutional neural networks, including architectures like DenseNet-121, have demonstrated strong performance in detecting a range of thoracic pathologies [1],[2]. However, successfully integrating such models into clinical workflows requires more than high classification accuracy. It demands robust mechanisms for managing uncertainty and ensuring patient safety.

H(pi)τiH(pi)≤τiH(pi)>τiFigure 1: Overview of the proposed selective chest X-ray classification framework. The multi-pathology classification model (a DenseNet-121 backbone) produces per-pathology probabilities; a confidence-estimation stage computes the binary entropy  of each, and quantile-based threshold calibration yields a per-class threshold . The decision module compares each pathology's entropy against its threshold, accepting confident findings () as a partial-pathology classification and deferring uncertain findings () for expert review. Each pathology is accepted or deferred independently, so acceptance is a partial-label output rather than a global clearance of the image.

One of the most critical risks in this domain is the occurrence of false negatives, where existing pathologies go undetected. Such failures can delay treatment and lead to serious consequences for patient outcomes. The problem is compounded by the nature of medical imaging, where visual data alone may not be sufficient for a definitive diagnosis. Identifying a condition may require image-based interpretation, contextual clinical information, laboratory results, or expert judgment. These challenges are amplified in multi-label classification settings, where multiple overlapping or subtle findings must be simultaneously considered.

In high-stakes medical imaging, a model that always produces a prediction, even when uncertain, can pose significant risks. Rejection mechanisms provide a safeguard by allowing the model to abstain on ambiguous cases, deferring them to expert review. This capability is especially critical in multi-pathology chest X-ray classification, where subtle or overlapping findings can lead to overconfident errors. Integrating rejection directly into the diagnostic workflow therefore enhances reliability and patient safety beyond what conventional classifiers can provide.

To address these concerns, we propose a chest X-ray classification framework built on a DenseNet-121 backbone and augmented with an entropy-based rejection mechanism that measures per-pathology uncertainty using binary entropy and abstains, independently for each pathology, when that uncertainty exceeds a calibrated threshold. Each pathology prediction is accepted or deferred on its own; acceptance yields a partial-label output and is not a global clearance of the image. This allows the model to selectively defer uncertain findings to human experts, reducing the risk of overconfident misclassifications.

Crucially, we implement a calibration procedure using a held-out portion of the training data to tune rejection thresholds. We explore two strategies: shared thresholds, which apply uniformly across all pathology classes, and class-specific thresholds, which are independently optimized per class to account for differences in prevalence and model confidence. This calibration-driven approach enables flexible and interpretable control over the trade-off between diagnostic performance and rejection coverage.

This study makes the following contributions to AI-assisted medical imaging: (1) an uncertainty-aware framework that integrates a DenseNet-121 backbone with a per-pathology entropy-based rejection mechanism for multi-pathology diagnosis; (2) a quantile-based calibration procedure that supports both global and class-specific thresholds, enabling fine-grained control over the trade-off between diagnostic performance and rejection coverage; (3) an evaluation based on risk-coverage analysis and balanced accuracy with bootstrap confidence intervals, together with a formal leave-one-dataset-out assessment of selective behaviour under domain shift across NIH ChestX-ray14, MIMIC-CXR, and the geographically independent PadChest cohort; and (4) the finding that pretrained chest X-ray backbones apply an operating-point normalization that deforms the per-pathology probability axis, so that predictive entropy must be computed on the raw model probability for the selective signal to transfer consistently across pathologies and datasets.

The remainder of this paper is organized as follows. Section 2 reviews related work on deep learning for chest X-ray diagnosis and selective classification. Section 3 describes the proposed methodology, including dataset construction, model design, and the rejection mechanism. Section 4 presents experimental results. Section 5 discusses the findings and their implications, and Section 6 concludes.

## 2. Prior work

Recent advances in deep learning have significantly improved automated chest X-ray interpretation, particularly in multi-label pathology classification. Deploying such models in clinical settings requires more than high predictive accuracy: it also demands robustness to uncertainty, transparency in decision-making, and mechanisms for risk-aware abstention. This section reviews three areas central to our work: deep learning for chest X-ray diagnosis, rejection mechanisms in machine learning, and selective prediction and uncertainty modeling in medical imaging.

### 2.1 Deep learning in medical imaging and chest X-ray diagnosis

Deep convolutional neural networks have become the method of choice for medical image analysis. Large public datasets and deep CNNs enable expert-level disease classification in chest radiography. Wang et al. [3] introduced the ChestX-ray8 dataset (approximately 112,000 frontal radiographs, 14 labels) and demonstrated multi-label classification and localization of common thoracic diseases. Rajpurkar et al. [4] trained a 121-layer DenseNet (CheXNet) on ChestX-ray14 and reported radiologist-level pneumonia detection. The CheXpert dataset [5] includes 224,316 chest radiographs with uncertainty labels and has been used to show that CNN models can reach radiologist-level performance on several pathologies. Lakhani and Sundaram [6] classified tuberculosis with an AUC near 0.99; Majkowska et al. [7] matched radiologist performance on five findings; and Park et al. [8] applied a self-evolving vision transformer with knowledge distillation. Cohen et al. [9] study balanced batch sampling across multiple chest X-ray datasets to improve out-of-distribution generalization for Cardiomegaly, Consolidation, Edema, and Effusion, using a DenseNet-121 with a leave-one-dataset-out strategy across ChestX-ray8 [3], CheXpert [5], MIMIC-CXR [10], and PadChest [11] [33]; we adopt the corresponding TorchXRayVision backbone [9] and augment it with a rejection mechanism. Cross-domain generalization in this setting is known to be fragile: Zech et al. [31] show that models can exploit hospital-specific confounders and degrade across acquisition sites, and Cohen et al. [33] attribute much of the cross-dataset gap to label shift rather than to a shift in the images themselves. Ovadia et al. [32] further show that predictive uncertainty degrades under dataset shift, motivating an explicit out-of-distribution evaluation of selective behaviour.

### 2.2 Rejection mechanisms in machine learning

The behavior of predictive models under uncertainty has been widely studied, with stability recognized as a key factor in the reliability of estimators. Prior work in statistical signal processing [12] and machine learning [13] shows that increased uncertainty leads to greater sensitivity to small perturbations, often indicating operation near a decision boundary. Such instability motivates analyzing perturbation responses as a diagnostic tool [14] and is frequently leveraged in active learning [15],[16], where uncertain samples are prioritized for annotation. Selective classification, the reject option, allows a model to abstain on uncertain inputs to reduce errors. Chow [17] formalized the error-reject trade-off, and Cortes et al. [18] developed frameworks for learning with a reject option. For deep networks, Geifman and El-Yaniv [19] introduced a risk-coverage framework that thresholds the maximal softmax response to guarantee a user-specified error rate, and proposed SelectiveNet [20], which optimizes the accuracy-coverage trade-off during training. Lakshminarayanan et al. [21] showed that deep ensembles provide high-quality uncertainty estimates. Fisch et al. [22] train a selector for calibrated selective classification; Cattelan and Silva [23] propose a logit normalization that improves confidence-based rejection; and Rabanser et al. [24] reject inputs whose predictions flip during training. The risk-coverage curve and its area introduced in this line of work [19] provide the threshold-free evaluation we adopt in Section 4.

### 2.3 Rejection methods in medical imaging and chest X-ray classification

Uncertainty-aware models can improve diagnostic safety by flagging ambiguous cases. Kompa et al. [25] emphasize that machine learning systems in healthcare should be able to say "I do not know" for high-uncertainty cases, and survey abstention methods. Faghani et al. [26] outline emerging trends in uncertainty quantification for medical imaging, noting that an uncertainty measure allows experts to re-evaluate doubtful cases. In chest radiography specifically, Whata et al. [27] developed a Bayesian CNN that quantifies predictive uncertainty for multi-class classification, and Das et al. [28] proposed a semi-supervised lesion-detection framework using confidence-guided pseudo-labelling. These works illustrate that rejection and uncertainty modelling are actively integrated into medical imaging pipelines to enhance decision support.

## 3. Methodology

Our framework for multi-label chest X-ray classification with selective prediction is illustrated in Figure 1. It integrates a deep convolutional neural network for pathology classification with a rejection mechanism that enables the model to abstain from uncertain predictions. The subsections below detail the dataset and partitioning strategies, the baseline model, the rejection mechanism, and the threshold calibration procedure.

### 3.1 Dataset

The classification task is formulated as multi-label, wherein each image can simultaneously exhibit one or more of four conditions: Cardiomegaly, Effusion, Edema, and Consolidation. In contrast to single-label classification, this setup accounts for the co-occurrence of multiple pathologies within a single image. To improve generalizability and account for variability in imaging protocols and patient populations, the dataset aggregates samples from multiple publicly available sources. Each image is labeled at the image level, and the labels are not mutually exclusive. The model produces a confidence score for each pathology, which is subsequently used by the rejection mechanism. Table 1 summarizes the class distribution of the final evaluation cohort (37,397 frontal images) used for all reported analyses, and Figure 2 shows representative frontal radiographs for the target conditions.

Table 1: Class distribution of the final evaluation cohort used for the selective-prediction analyses, across the three sources (NIH ChestX-ray14, MIMIC-CXR, and PadChest; N = 37,397 frontal images), reported as positive-label counts. This cohort is the harmonized, patient-level-split set on which all reported results are computed. The data is imbalanced, with Effusion and Cardiomegaly more prevalent than the rarer Edema and Consolidation, and a subset of images carry multiple positive labels.

| Category | Positive count | Percentage (%) |
| --- | --- | --- |
| Cardiomegaly | 3447 | 9.2 |
| Effusion | 4396 | 11.8 |
| Edema | 1407 | 3.8 |
| Consolidation | 1125 | 3.0 |
| Multi-Pathology | 1885 | 5.0 |

Figure 2: Representative frontal chest radiographs spanning the range of thoracic findings in the source datasets. The four conditions studied in this work, Cardiomegaly, Effusion, Edema, and Consolidation, are a subset of these findings; the selective mechanism makes an independent accept-or-defer decision for each of the four.

### 3.2 Data partition for validation and rejection tuning

The dataset aggregates chest X-ray images from MIMIC-CXR, NIH ChestX-ray14, and PadChest. Each source uses slightly different imaging protocols, labelling criteria, and patient populations, introducing domain variability. We employ two partitioning strategies. In intra-source splitting, each dataset is split independently into 80% for training and 20% for validation, with all partitions performed at the patient level so that images from the same patient are assigned exclusively to one subset, preventing leakage and reflecting deployment to previously unseen individuals. In inter-source splitting, one or two complete source datasets are used for training while another source is held out entirely for evaluation, creating an out-of-distribution test scenario. The intra-source approach supports stable tuning under controlled conditions, while the inter-source setup tests robustness to domain shift.

### 3.3 Baseline model

DenseNet-121 [29] is a convolutional neural network characterized by dense connectivity, wherein each layer receives the feature maps of all preceding layers within a dense block. This design enhances information flow, alleviates the vanishing-gradient problem, and promotes feature reuse. The architecture consists of an initial convolution and pooling layer followed by four dense blocks separated by three transition layers, each dense block following a BN, ReLU, Conv(3x3) pattern, with a growth rate that governs the progressive expansion of feature dimensionality.

The baseline employs a standard DenseNet-121 backbone pre-trained on large-scale chest X-ray datasets, including CheXpert and NIH ChestX-ray14 [4],[5]. We employ a fine-tuning strategy with multi-domain balanced sampling, constructing each mini-batch to include an equal number of samples from each domain to encourage domain-invariant features [9]. The model performs multi-label classification by outputting a probability for each target pathology. Training uses the AdamW optimizer (β1=0.9, β2=0.999, learning rate 10−4, batch size 64) for up to 30 epochs, with a ReduceLROnPlateau schedule (factor 0.2, patience 4) and square-root inverse-frequency class weighting within a weighted binary cross-entropy loss to mitigate class imbalance while maintaining calibrated probability estimates. Input images are converted to single-channel grayscale, center-cropped, resized to 224x224, and intensity-normalized to the [-1024, 1024] range expected by the backbone. Calibration and evaluation splits are patient-level, and the random seed is fixed (1337) throughout; code and frozen per-image predictions are released for full reproducibility. The primary evaluation uses the released TorchXRayVision DenseNet-121 [9], whose published checkpoint was trained on a pooled corpus that includes NIH ChestX-ray14, MIMIC-CXR, CheXpert, and PadChest; consequently MIMIC-CXR and PadChest are in-distribution for this backbone, and the replication results in Section 4.1 establish that the selective benefit holds across sources rather than constituting an out-of-distribution test. The out-of-distribution claims (Section 4.3) instead rest on single-source leave-one-dataset-out models and on the independently trained weighted-BCE DenseNet-121, which follows the procedure above and is evaluated on a source it never trained on. For the primary evaluation the released weights are used and the rejection mechanism is applied post hoc to frozen predictions; the training hyperparameters above describe the independently trained weighted-BCE replication model, not the released backbone.

### 3.4 Rejection mechanism

Rejection is defined independently for each pathology. For the predicted probability pi of pathology i, the binary entropy is

H(pi)=− pilog2pi+(1−pi)log2(1−pi) ,

which is maximal at the decision boundary and minimal as pi→0 or pi→1. A pathology prediction is rejected when its entropy exceeds a calibrated threshold τi,

reject class i⇔H(pi)>τi.

The accept-defer boundary is therefore H(pi)=τi, where τi is the calibrated entropy threshold for pathology i. This rejection threshold is distinct from the probability threshold used to convert pi into a binary label; the latter sets the decision boundary, while τi sets how much uncertainty around that boundary is tolerated before a finding is deferred.

Selective prediction is applied independently for each pathology: a finding is deferred when the model cannot generate a sufficiently reliable prediction for that condition. Acceptance corresponds to a partial-label output and is not a global clearance of the image, since other findings for the same image may remain uncertain. Rejection rates are therefore computed and reported per pathology, aligning with a clinical decision-support workflow in which uncertain findings are deferred for expert review while confident predictions remain available.

An equivalent view of this rule is interval-based: a class is treated as confident when a calibrated interval around its probability lies entirely on one side of the decision boundary. On a deterministic backbone this interval rule is a monotone function of the distance of the probability from the boundary, so it induces the same selective ordering as entropy thresholding; we therefore adopt the single entropy-based rule. Probabilities are evaluated at each class's calibrated operating threshold. Thresholds are selected globally or per class on a held-out calibration subset.

Predictive entropy is computed on the backbone's raw per-pathology probability (the sigmoid output), with the operating-point normalization that the pretrained checkpoints apply by default disabled. That normalization rescales each pathology's output around a fixed training-source operating threshold, deforming the probability axis differently per pathology and per source, which distorts the entropy ranking used for deferral; the raw probability provides a consistent uncertainty axis across pathologies and datasets.

### 3.5 Threshold calibration

We adopt a quantile-based calibration strategy to determine effective thresholds. Calibration is performed on a dedicated subset disjoint from both the training and evaluation data, minimizing optimistic bias from information leakage. Entropy thresholds are estimated from the distribution of entropy values computed over correctly classified calibration samples, identifying confidence regions historically associated with reliable predictions. For each pathology, we explore a range of candidate percentiles (from 0.70 to 0.99) and compute either a class-specific or a global threshold corresponding to the selected quantile. A threshold-sensitivity analysis (Figure 7) reports the selective metric and rejection rate across calibration percentiles from 0.70 to 0.99, so that the operating point is an interpretable, tunable design choice rather than a fixed property of the model. Although results are reported at fixed thresholds for consistency, the mechanism outputs a continuous confidence score, allowing thresholds to be tuned post hoc to a clinical risk tolerance.

## 4. Results

Calibration and evaluation use disjoint patient-level splits. Results are reported on 11,212 NIH ChestX-ray14 frontal images, 8,185 MIMIC-CXR frontal images, and 18,000 PadChest frontal images. The backbone attains per-pathology discrimination in the range of AUC 0.77 to 0.94 across the three evaluation sources. We characterize selective behaviour with two complementary, threshold-aware measures: balanced accuracy at a calibrated operating point, and the risk-coverage curve. Both are reported with 1000-sample non-parametric bootstrap confidence intervals; these intervals are robust to patient-level clustering of the resampling, which shifts every interval bound by at most 0.002 and changes no support decision.

### 4.1 Selective prediction improves balanced accuracy

Predictive entropy is informative about correctness: incorrect predictions concentrate at high entropy and correct predictions at low entropy (Figure 3), the signal that selective deferral exploits.

Figure 3: Distribution of predictive entropy for correct versus incorrect predictions, per pathology (rows) and dataset (columns: NIH ChestX-ray14, MIMIC-CXR, PadChest). For every pathology on every dataset, correct predictions concentrate at low entropy and incorrect predictions at high entropy, which is what makes uncertainty-based deferral effective.

At a calibrated operating point, deferring the most uncertain predictions improves balanced accuracy for every target pathology. Table 2 reports the balanced accuracy at full coverage and at a coverage of 0.70 (a rejection rate of 0.30) on NIH ChestX-ray14, with bootstrap confidence intervals on the change. All four improvements are supported (95% confidence intervals excluding zero), and the effect replicates on MIMIC-CXR and on an independently trained model; Figure 4 traces the balanced accuracy across coverage for each pathology.

Table 2: Selective balanced accuracy at a calibrated operating point (NIH ChestX-ray14, rejection rate 0.30). All four gains have 95% bootstrap confidence intervals excluding zero.

| Pathology | Balanced accuracy (full coverage) | Balanced accuracy at 0.70 coverage | Change (95% CI) |
| --- | --- | --- | --- |
| Cardiomegaly | 0.757 | 0.844 | +0.087 [0.058, 0.115] |
| Effusion | 0.780 | 0.855 | +0.075 [0.064, 0.087] |
| Edema | 0.754 | 0.839 | +0.085 [0.052, 0.123] |
| Consolidation | 0.730 | 0.797 | +0.067 [0.049, 0.086] |

Figure 4: Selective balanced accuracy as a function of coverage for each pathology. Balanced accuracy increases as the most uncertain predictions are deferred.

On MIMIC-CXR (8,185 frontal images), the same calibrated-threshold analysis yields supported balanced-accuracy gains for all four pathologies (Cardiomegaly +0.071 [0.059, 0.086], Effusion +0.086 [0.073, 0.097], Edema +0.072 [0.058, 0.087], Consolidation +0.034 [0.014, 0.054]; 95% confidence intervals excluding zero), and an independently trained weighted-BCE DenseNet-121 (Section 3.3; validation macro-AUC 0.85) reproduces the improvement on all four pathologies both on the held-out NIH test set (Cardiomegaly +0.074 [0.049, 0.098], Effusion +0.062 [0.052, 0.073], Edema +0.060 [0.022, 0.101], Consolidation +0.044 [0.012, 0.073]) and cross-dataset on MIMIC-CXR, a source it never trained on (Cardiomegaly +0.055 [0.042, 0.066], Effusion +0.087 [0.075, 0.098], Edema +0.073 [0.057, 0.090], Consolidation +0.067 [0.036, 0.100]); all gains have 95% confidence intervals excluding zero, confirming the effect is specific to neither the backbone nor the dataset. On PadChest, where the backbone attains strong discrimination (Cardiomegaly AUC 0.91, Effusion 0.94, Consolidation 0.87), deferral improves balanced accuracy for all four pathologies (Cardiomegaly +0.080 [0.070, 0.091], Effusion +0.056 [0.043, 0.070], Edema +0.075 [0.058, 0.096], Consolidation +0.076 [0.051, 0.108]; 95% confidence intervals excluding zero), confirming the benefit on a third source; the corresponding out-of-distribution test, in which a single-source model is applied to PadChest, is reported in Section 4.3.

A non-parametric bootstrap confirms these gains are not resampling artefacts: across 1,000 resamples the balanced accuracy with rejection exceeds the baseline for every pathology on all three evaluation sets (Figure 5).

Figure 5: Bootstrap validation of the balanced-accuracy change before and after rejection. Boxplots show the distribution of balanced accuracy over 1,000 bootstrap resamples at full coverage (baseline, blue) and at a coverage of 0.70 (with rejection, orange), per pathology, on the NIH ChestX-ray14, MIMIC-CXR, and PadChest evaluation sets. The with-rejection distribution exceeds the baseline for every pathology on all three datasets.

To confirm that these gains come from the uncertainty signal rather than from evaluating on a smaller retained set, entropy-based deferral is compared against a random-rejection control and against alternative confidence scores on the same NIH evaluation split. At a coverage of 0.70, random rejection leaves balanced accuracy essentially unchanged (within 0.001 of zero for all four pathologies), whereas entropy-based deferral recovers the gains of Table 2. Margin and maximum-probability rejection are rank-identical to entropy on a deterministic backbone and induce the same selective ordering, so entropy is reported as their representative. A deep-ensemble disagreement signal computed across the five public TorchXRayVision checkpoints does not exceed single-pass entropy at this operating point, so the single forward-pass rule is retained.

### 4.2 Risk-coverage analysis

Figure 6 presents the risk-coverage curve for each pathology, summarized by the area under the curve. As coverage decreases, the model retains its more confident predictions, and the curves quantify the accuracy-coverage trade-off that a deployment can select according to a target risk tolerance.

Figure 6: Risk-coverage curves per pathology. The retained-prediction performance is traced against coverage as the entropy threshold varies. Area under the risk-coverage curve over the plotted coverage range (AURC; lower is better): Cardiomegaly 0.206, Effusion 0.180, Edema 0.207, Consolidation 0.239.

### 4.3 Out-of-distribution evaluation

To assess behaviour under domain shift, we evaluate a single-source model on an unseen source in a leave-one-dataset-out protocol (Table 3). The discrimination of a model trained on NIH ChestX-ray14 decreases when it is applied to MIMIC-CXR, quantifying the domain-shift penalty that the inter-source setting is designed to expose and against which selective prediction provides a safeguard by deferring uncertain cases.

Table 3: Leave-one-dataset-out evaluation (model trained on NIH ChestX-ray14, evaluated on MIMIC-CXR). Discrimination decreases on the unseen source, quantifying the domain-shift penalty.

| Pathology | In-domain AUC (NIH) | Out-of-distribution AUC (MIMIC) |
| --- | --- | --- |
| Cardiomegaly | 0.864 | 0.687 |
| Effusion | 0.818 | 0.793 |
| Edema | 0.867 | 0.763 |
| Consolidation | 0.793 | 0.694 |

Domain shift lowers discrimination, but the selective mechanism remains effective on the shifted distribution: deferring the most uncertain predictions improves retained balanced accuracy on the unseen source. Table 4 reports this for the NIH-trained model evaluated on MIMIC-CXR, where all four gains at a coverage of 0.70 have 95% bootstrap confidence intervals excluding zero. The same pattern holds in the reverse direction (a MIMIC-trained model evaluated on NIH ChestX-ray14): deferral improves retained balanced accuracy for all four pathologies (Cardiomegaly +0.098 [0.004, 0.204], Effusion +0.081 [0.050, 0.109], Edema +0.138 [0.055, 0.225], Consolidation +0.063 [0.014, 0.109]). The robustness extends to a geographically independent source: applying the NIH-trained model to PadChest, a Spanish cohort never seen during training, deferral improves retained balanced accuracy for all four pathologies (Cardiomegaly +0.078 [0.065, 0.090], Effusion +0.074 [0.058, 0.088], Edema +0.045 [0.025, 0.067], Consolidation +0.059 [0.007, 0.107]), all with 95% confidence intervals excluding zero. Selective prediction therefore provides a safeguard precisely where it is most needed, on out-of-distribution inputs.

Table 4: Selective balanced accuracy under domain shift (NIH-trained model evaluated on MIMIC-CXR), at full coverage and at a coverage of 0.70. All four gains have 95% bootstrap confidence intervals excluding zero.

| Pathology | Balanced accuracy (full coverage) | Balanced accuracy at 0.70 coverage | Change (95% CI) |
| --- | --- | --- | --- |
| Cardiomegaly | 0.632 | 0.676 | +0.044 [0.033, 0.058] |
| Effusion | 0.716 | 0.787 | +0.071 [0.059, 0.081] |
| Edema | 0.673 | 0.695 | +0.022 [0.010, 0.035] |
| Consolidation | 0.619 | 0.668 | +0.049 [0.022, 0.076] |

### 4.4 Threshold sensitivity

Figure 7 traces the rejection rate and retained performance across calibration percentiles. Selective performance varies smoothly with the calibration percentile, so the accuracy-coverage trade-off is a controllable design parameter that can be set to a clinical risk tolerance. At a fixed calibration percentile the induced rejection rate is similar across pathologies (0.37 to 0.43 at the 0.70 percentile), reflecting the consistent raw-probability uncertainty axis on which entropy is computed.

Figure 7: Threshold sensitivity. Rejection rate and retained performance as a function of the calibration percentile.

### 4.5 Decision-region behaviour and interpretability

The decision regions of accepted and deferred predictions can be examined spatially with class-activation mapping (Figure 8): for confident accepted predictions the activation concentrates over the lung fields, whereas borderline deferred predictions show more diffuse activation, consistent with the uncertainty-driven abstention used throughout this work.

Figure 8: Representative Grad-CAM overlays for confident accepted predictions (top) and borderline deferred predictions (bottom). Accepted predictions show activation concentrated over the lung fields, whereas deferred predictions show more diffuse activation.

## 5. Discussion

The results show that selective prediction provides meaningful control over the trade-off between predictive confidence and coverage. At a calibrated operating point, deferring the most uncertain per-pathology predictions improves balanced accuracy for all four conditions on NIH ChestX-ray14 and MIMIC-CXR, on an independently trained model (including when applied cross-dataset), and under leave-one-dataset-out domain shift in both transfer directions, with every target pathology supported on all three datasets; the risk-coverage curves give a complete, threshold-aware view of the accuracy-coverage trade-off. The improvement is most pronounced for conditions where the model is less certain, consistent with rejection concentrating on ambiguous and low-prevalence findings such as Edema.

A leave-one-dataset-out evaluation quantifies the domain-shift penalty and shows that deferring uncertain cases improves retained balanced accuracy on the unseen source in both transfer directions, so selective prediction is most valuable precisely on the out-of-distribution inputs that most threaten reliability. The operating point used throughout defers 30% of per-pathology predictions to recover balanced-accuracy gains of 0.07 to 0.09 on NIH ChestX-ray14; because the rejection rate is a tunable, per-pathology quantile (Figure 7), a deployment can trade reviewer workload against retained accuracy to match its capacity, deferring fewer cases where expert review is scarce and more where the cost of an error is high. A random-rejection control that defers the same fraction yields no such gain (Section 4.1), so the benefit reflects the uncertainty ranking rather than the reduced caseload.

These findings extend the selective-classification framework, which formalizes the error-reject trade-off [17] and the risk-coverage view [19], to the multi-label medical-imaging setting, and they isolate a point specific to it: at a calibrated per-pathology operating point the reject option raises balanced accuracy even where it cannot raise a rank-based measure such as AUC, because abstention removes the near-boundary cases on which ranking depends. Whereas deep ensembles derive uncertainty from multiple forward passes [21], the single-pass entropy rule attains comparable selective behaviour here (Section 4.1), and the gains persist under the dataset shift known to degrade uncertainty estimates [32] and cross-site generalization [31]. Together these findings support uncertainty-aware selective prediction as a route to safer, more trustworthy AI-assisted diagnosis.

The cross-source behaviour is consistent with the documented fragility of chest X-ray models across acquisition sites, where networks exploit hospital- and protocol-specific confounders [31], and with label-definition shift between sources rather than a shift in the images themselves [33]. The datasets differ in scanner and protocol, patient population (PadChest is a Spanish outpatient cohort, MIMIC-CXR an intensive-care population), and in projection mix. For Cardiomegaly this projection mix is decisive: the cardiothoracic ratio that defines cardiomegaly is geometrically magnified on anteroposterior and portable acquisitions (the heart lies farther from the detector), so a model trained on a predominantly posteroanterior cohort over-calls cardiomegaly on anteroposterior-heavy sources. Stratifying by view makes this explicit: within the posteroanterior stratum the NIH-trained model transfers to PadChest with AUC 0.88 and deferral still improves balanced accuracy (+0.10 at a coverage of 0.70), whereas on the anteroposterior and portable subset discrimination falls to AUC 0.66. This is a geometric property of acquisition, not a pixel artifact that intensity normalization or contrast windowing can remove; selective prediction is the appropriate response, abstaining on the projections where the cue is unreliable.

### 5.1 Limitations

The validated regime has five boundary conditions. First, the evaluation covers four thoracic findings on frontal views; additional findings and lateral views remain for future work. Second, the primary backbone is trained on a corpus that includes the MIMIC-CXR and PadChest evaluation sources (Section 3.3), so the cross-source generalization claims rest on the single-source leave-one-dataset-out models and the independently trained weighted-BCE model rather than on the primary backbone alone. Third, results are reported for the DenseNet-121 family; transfer to other backbones is established only through the independently trained model. Fourth, the labels are the public datasets' image-level annotations, which carry known noise, and the abstention benefit is measured against these labels rather than a prospective radiologist-adjudicated reference; a reader study quantifying the clinical value of the deferred-case workflow is a natural next step. Finally, subgroup fairness of the selective mechanism, that is, whether deferral rates and retained accuracy are equitable across patient subgroups [30], is not assessed here and is left for future work.

## 6. Conclusions and future directions

We presented a framework for multi-label chest X-ray diagnosis that integrates a DenseNet-121 backbone with a per-pathology entropy-based rejection mechanism and a quantile-based calibration of global or class-specific thresholds. Evaluated with risk-coverage analysis and balanced accuracy under bootstrap confidence intervals, and with a formal leave-one-dataset-out assessment, selective prediction improves the reliability of accepted predictions and provides interpretable control over the accuracy-coverage trade-off. The validated regime and its boundary conditions are detailed in Section 5.1. Future work will incorporate complementary clinical information, extend the approach to additional disease classes, and explore Bayesian and ensemble uncertainty estimators. Code and frozen per-image predictions are released to support reproduction.

## Data Availability

All datasets used in this study are publicly available. CheXpert: stanfordmlgroup.github.io/competitions/chexpert/; MIMIC-CXR-JPG: physionet.org/content/mimic-cxr-jpg/2.1.0/; NIH ChestX-ray14: nihcc.app.box.com/v/ChestXray-NIHCC; PadChest: bimcv.cipf.es/bimcv-projects/padchest/.

## Code Availability

Code, configuration files, and frozen per-image predictions used to reproduce every table and figure in this paper are available at github.com/ApartsinProjects/cxr-selective-rejection-code.

## Funding

This research received no specific grant from any funding agency in the public, commercial, or not-for-profit sectors.

References

- 1. Sufian, M. A., et al. (2024). AI-Driven Thoracic X-ray Diagnostics: Transformative Transfer Learning for Clinical Validation in Pulmonary Radiography. Journal of Personalized Medicine, 14(8), 856.

- 2. Arulananth, T. S., et al. (2024). Classification of pediatric pneumonia using modified DenseNet-121 deep-learning model. IEEE Access.

- 3. Wang, X., Peng, Y., Lu, L., Lu, Z., Bagheri, M., Summers, R. M. (2017). ChestX-ray8: Hospital-scale chest X-ray database and benchmarks. CVPR 2017, 3462-3471.

- 4. Rajpurkar, P., et al. (2017). CheXNet: Radiologist-level pneumonia detection on chest X-rays with deep learning. arXiv:1711.05225.

- 5. Irvin, J., et al. (2019). CheXpert: A large chest radiograph dataset with uncertainty labels and expert comparison. AAAI, 33(1), 590-597.

- 6. Lakhani, P., Sundaram, B. (2017). Deep learning at chest radiography: Automated classification of pulmonary tuberculosis. Radiology, 284(2), 574-582.

- 7. Majkowska, A. E., et al. (2020). Chest radiograph interpretation with deep learning models. Radiology, 294(2), 421-431.

- 8. Park, S., et al. (2022). Self-evolving vision transformer for chest X-ray diagnosis through knowledge distillation. Nature Communications, 13, 3848.

- 9. Cohen, J. P., Viviano, J., Morrison, P., Brooks, R., Hashir, M., Bertrand, H. (2020). TorchXRayVision: A library of chest X-ray datasets and models. arXiv:2111.00595.

- 10. Johnson, A. E., et al. (2019). MIMIC-CXR-JPG, a large publicly available database of labeled chest radiographs. arXiv:1901.07042.

- 11. Bustos, A., Pertusa, A., Salinas, J. M., de la Iglesia-Vaya, M. (2020). PadChest: A large chest X-ray image dataset with multi-label annotated reports. Medical Image Analysis, 66, 101797.

- 12. Apartsin, A., Cooper, L. N., Intrator, N. (2012). Semi-coherent time of arrival estimation using regression. JASA, 132(2), 832-837.

- 13. Goodfellow, I. J., Shlens, J., Szegedy, C. (2014). Explaining and harnessing adversarial examples. arXiv:1412.6572.

- 14. Lee, J., AlRegib, G. (2020). Gradients as a measure of uncertainty in neural networks. ICIP 2020, 2416-2420.

- 15. Rokach, L., Aperstein, Y., Akselrod-Ballin, A. (2025). Deep active learning framework for chest-abdominal CT scans segmentation. Expert Systems with Applications, 263, 125522.

- 16. Levy, D., et al. (2026). From Cold Start to Active Learning: Embedding-Based Scan Selection for Medical Image Segmentation. arXiv:2601.18532.

- 17. Chow, C. (1970). On optimum recognition error and reject tradeoff. IEEE Transactions on Information Theory, 16(1), 41-46.

- 18. Cortes, C., DeSalvo, G., Mohri, M. (2016). Learning with rejection. Journal of Machine Learning Research, 17(216), 1-40.

- 19. Geifman, Y., El-Yaniv, R. (2017). Selective classification for deep neural networks. NeurIPS 30, 4878-4887.

- 20. Geifman, Y., El-Yaniv, R. (2019). SelectiveNet: A deep neural network with an integrated reject option. ICML 2019, 2151-2160.

- 21. Lakshminarayanan, B., Pritzel, A., Blundell, C. (2017). Simple and scalable predictive uncertainty estimation using deep ensembles. NeurIPS 30, 6405-6416.

- 22. Fisch, A., Jaakkola, T. S., Barzilay, R. (2022). Calibrated selective classification. Transactions on Machine Learning Research.

- 23. Cattelan, L., Silva, R. (2024). How to fix a broken confidence estimator: evaluating post-hoc methods for selective classification with deep neural networks. UAI 2024. arXiv:2305.15508.

- 24. Rabanser, S., Thudi, A., Hamidieh, K., Dziedzic, A., Papernot, N. (2022). Selective classification via neural network training dynamics. UAI 2022. arXiv:2205.13532.

- 25. Kompa, B., Snoek, J., Beam, A. L. (2021). Second opinion needed: Communicating uncertainty in medical machine learning. NPJ Digital Medicine, 4, 4.

- 26. Faghani, S., et al. (2023). Quantifying uncertainty in deep learning of radiologic images. Radiology, 308(2).

- 27. Whata, A., et al. (2024). Uncertainty quantification in multi-class image classification using chest X-ray images of COVID-19 and pneumonia. Frontiers in Artificial Intelligence, 7, 1410841.

- 28. Das, A., et al. (2024). Confidence-guided semi-supervised learning for generalized lesion localization in X-ray images. MICCAI 2024, LNCS 15001, 242-252.

- 29. Huang, G., Liu, Z., Van Der Maaten, L., Weinberger, K. Q. (2017). Densely connected convolutional networks. CVPR 2017, 4700-4708.

- 30. Seyyed-Kalantari, L., Liu, G., McDermott, M., Chen, I. Y., Ghassemi, M. (2020). CheXclusion: Fairness gaps in deep chest X-ray classifiers. Pacific Symposium on Biocomputing 2021, 232-243.

- 31. Zech, J. R., Badgeley, M. A., Liu, M., Costa, A. B., Titano, J. J., Oermann, E. K. (2018). Variable generalization performance of a deep learning model to detect pneumonia in chest radiographs. PLOS Medicine, 15(11), e1002683.

- 32. Ovadia, Y., et al. (2019). Can you trust your model's uncertainty? Evaluating predictive uncertainty under dataset shift. NeurIPS 32.

- 33. Cohen, J. P., Hashir, M., Brooks, R., Bertrand, H. (2020). On the limits of cross-domain generalization in automated X-ray prediction. Medical Imaging with Deep Learning (MIDL), PMLR 121, 136-155. arXiv:2002.02497.